## Slide 1: Design Updates: September 15, 2014
Most of the components have arrived; remaining items:
  PV system
    solar panels
    charge controller
    cables, switches/breakers
    Mounting hardware
  Trailer (for PV array + batteries & transporting system)
  Pump (discussion follows)
  (?) Pressure gage (following pump discussion)
Major decision:
  retain single tube with subsampling into canisters (high power consumption)
  or sample directly from atmosphere into canisters (low power consumption)

## Slide 2: Under current design:
To ensure turbulent flow in single tube, minimum flow rate (L/min):
Tube size | @ Sea level | Pullman | Denver
Standard ¼” | 11.2 | 12.5 | 13.6
Thick-wall ¼” | 9.1 | 10 | 11
Standard 1/8” | 2.3 | 2.5 | 2.7
Target pump operating conditions: 2.5 L/min at 400 mbar
Based on.. | @ Sea level | Pullman | Denver
Pressure drops | 484 | 437 | 350
Canister fill safety margin | 434 | 395 | 357
To ensure choked flow across orifice, max. downstream pressure (mbar):

## Slide 3: Pump candidates:
Mnfctr | Model | Power | Type | Notes
KNF | NPK09DC | 62W | Diaphragm | Barely meets operating conditions
KNF | NPK 018…DC | 130W | Diaphragm | Moderate operating margin
Gast | 15D1150-101-1006 | 13W | Diaphragm | Barely meets operating range; power specs are suspiciously optimistic
Gast | 22D1180-201-1002 | 96W | Diaphragm | Narrow operating margin
Gast | 1531-107B-G617 | 170W | Rotary vane | Very wide operating margin; power specs missing, reported value is from pump with similar characteristics
Gast | LOA-101-JR | ~120W | Rocking piston | Moderate operating margin Power specs are 11-15V @ 9A Finicky: temp range (5-40 ºC), RH range (20-80%)
Gast | ROA-V210-JH / ROA-P201-JH | 170W | Rocking piston | Wide operating margin; Finicky: temp range (5-40 ºC), RH range (20-80%)
Air Dimensions | J161-xx-HH0 | 29W | Diaphragm | Barely almost hardly meets operating conditions
Air Dimensions | 201-xx-JH0 | 460W | Diaphragm | Barely meets operating conditions; power specs missing, reported value is from 222 series (dual head version of 221-)
Air Dimensions | 221-xx-JH0 | 460W | Diaphragm | 
Choice assumed in ‘Consumption’ spreadsheet

## Slide 4
With indicated pump choice, the solar array would need to be about 1000W to provide 365/24/7 operation. However:
pump indicated is probably insufficient to achieve turbulent flow
larger pump raises the req'd array size EX: the other KNF model listed is 130W -> raises array to 1800W
Can we lower system power consumption to minimize solar array size?
Sure, mitigation strategies include:
  Seasonal instead of year-round operation
  Not operating at night
  Bye-days
  Deploy at lower latitudes (this is destined for CA, right?)
Other ideas:
  Modify sampling configuration (discussion begins next slide)
  Automatic generator back-up (yuck)
  Small wind-turbine generator (except for possibly disrupting wind field?)

## Slide 5
move orifices/3-way valves to sampling location (diagram)
use small dual-head pump to draw air across orifices such as Brailsford TD-4X2N (pictured)
  with heads in series, it produces sufficient vacuum for choked flow (measured @ -19 in Hg ≈ -650 mb)
  small power consumption (0.4A @ 9-14Vdc) eliminates need to power pump directly off PV array batteries
One alternative:
to pump
(normally open)
to canister
(normally closed)
orifice (common)
negligible-length
1/8" leader tube
P
Buy another pressure sensor?:
To detect pump/diaphragm failure, and to ensure sufficient vacuum is present for choked flow, add a pressure sensor between each pump & 3-way valve.
One electronic 0-15 psia gage (Omega PX309-015G5V) is available but one more would be required at an estimated $240.

## Slide 6
Relocating orifice/3-way valve assembly to near sonic sampling volume introduces volume between canister & orifice. However, this volume is negligible.
Even under the following assumptions, volume of tube from 3-way valve to canister is less than 1% of actual volume in canister:
10% safety margin (canister filled to 42.8% of capacity or about 14 std L)
30 ft of standard 1/4" tubing
With 20 ft of 1/8" standard, volume in tubing is 0.02% of final sample volume.
Tube size | 10 ft | 15ft | 20 ft | 25 ft | 30 ft
Standard 1/4" | 37 | 56 | 75 | 94 | 113
Thick-wall 1/4" | 25 | 37 | 49 | 61 | 74
Standard 1/8" | 1.5 | 2.2 | 3.0 | 3.7 | 4.5
Volume (mL) as a function of tube size, length:

## Slide 7
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
Sampling
Canister
Closed
Open
P1
Open
Smart REA Sampler
Collection Mode
Identical systems for ‘up’ and ‘down’ sampling.
Critical orifice mounts in common side of solenoid valve.
Pressure gauge #1 needs to work across wide range… something like <0.1 psia to 30 psia.
Pressure gauge #2 operates under vacuum, so general purpose 0-15 psia is satisfactory.
Manual valves must be leak tight in vacuum, and solenoid must seal under same conditions.
Tubing downstream of solenoid is 1/8" Teflon.
Fittings from valve(s) to canister are stainless steel.
Filter probably needed upstream of pump (not shown).
Switching
3
3
Air parcel
P2
To
Pump

## Slide 8
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
Sampling
Canister
Open
Open
P1
Closed
Smart REA Sampler
Pressurization Mode
After sampling period, canister is filled to ~30 psi with UHP nitrogen.
Pressure gauge records pressure change so that the dilution factor is precisely known.
From
Regulated
Gas Cylinder
Closed
to Pump
3
To
Pump
3
P2
Open to Pump

## Slide 9
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
Sampling
Canister
Open
Open
P1
Closed
Smart REA Sampler
Transfer Mode
Pressurized main canister is used to fill several smaller canisters/cartridges.
Cartridges are attached in parallel so that all are filled to same pressure.
Cartridges can be then transferred to laboratory for analysis.
Pressure gauge records pressure change so that the cartridge pressures are known.
Cartridge design still TBD.
To Smaller
Evacuated
Cartridges
3
Closed
to Pump
Open to Pump
To
Pump
Closed
to Pump
3
To
Pump
P2

## Slide 10
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
Sampling
Canister
Open
Open
P1
Open
Smart REA Sampler
Evacuation Mode
Vacuum pump must be able to evacuate system to near vacuum (< 0.1 psi = 7 mb).
Solenoid also evacuated to minimize contamination.
To Suitable
Vacuum
Pump
3
Closed
to Pump
Open to Pump
To
Pump
Closed
to Pump
3
To
Pump
P2

## Slide 11: Design Updates: June 11, 2013
Contacted Tom Jobson & Claudia Toro about borrowing canisters. Claudia will work with Korey to get trained up on how to evacuate/fill them.
New design for the switching mechanism. Swapped critical orifice with switching valve and changed switching valve to three-way valve. Other arm of three-way valve will be tied to vacuum pump.
  Flow will be passing through the critical orifice at all time, addressing concern about transient flows through the orifice.
  Dead volume less of a concern now. Instead, we just need to have a very good understanding of the expected travel time from the tower top to the three way valve. Note Patrick’s concern below.
Patrick O’Keefe asked how we “anticipate handling the time delay between the measurement of an up/down eddy at the sonic and the switching of the appropriate valve down at the canister? Will the flow be regulated well-enough to use a static pre-calculated delay time? Seems like the switching should occur at the sonic to avoid that ambiguity.”

## Slide 12
S
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Solenoid Valve
From
Tower
To
Pump
Sampling
Canister
Closed
Open
P
S
Open
Switching
Smart REA Sampler
Collection Mode
Identical systems for ‘up’ and ‘down’ sampling.
Need to minimize dead volume between solenoid and critical orifice.
Pressure gauge needs to work across wide range… something like <0.1 psi to 30 psi.
Manual valves must be leak tight in vacuum, and solenoid must seal under same conditions.
All tubing & fittings downstream of solenoid are stainless steel.
Filter needed upstream of solenoid (not shown).
OLD DESIGN 1

## Slide 13
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
From
Tower
To
Pump
Sampling
Canister
Closed
Open
P
Open
Smart REA Sampler
Collection Mode
Identical systems for ‘up’ and ‘down’ sampling.
Need to minimize dead volume between solenoid and critical orifice.
Pressure gauge needs to work across wide range… something like <0.1 psi to 30 psi.
Manual valves must be leak tight in vacuum, and solenoid must seal under same conditions.
All tubing & fittings downstream of solenoid are stainless steel.
Filter needed upstream of solenoid (not shown).
Switching
3
To
Pump
3
OLD DESIGN 2