## Slide 1: Design Updates: June 11, 2013
Contacted Tom Jobson & Claudia Toro about borrowing canisters. Claudia will work with Korey to get trained up on how to evacuate/fill them.
New design for the switching mechanism. Swapped critical orifice with switching valve and changed switching valve to three-way valve. Other arm of three-way valve will be tied to vacuum pump.
  Flow will be passing through the critical orifice at all time, addressing concern about transient flows through the orifice.
  Dead volume less of a concern now. Instead, we just need to have a very good understanding of the expected travel time from the tower top to the three way valve. Note Patrick’s concern below.
Patrick O’Keefe asked how we “anticipate handling the time delay between the measurement of an up/down eddy at the sonic and the switching of the appropriate valve down at the canister? Will the flow be regulated well-enough to use a static pre-calculated delay time? Seems like the switching should occur at the sonic to avoid that ambiguity.”

## Slide 2
S
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Solenoid Valve
From
Tower
To
Pump
Sampling
Canister
Closed
Open
P
S
Open
Switching
Smart REA Sampler
Collection Mode
Identical systems for ‘up’ and ‘down’ sampling.
Need to minimize dead volume between solenoid and critical orifice.
Pressure gauge needs to work across wide range… something like <0.1 psi to 30 psi.
Manual valves must be leak tight in vacuum, and solenoid must seal under same conditions.
All tubing & fittings downstream of solenoid are stainless steel.
Filter needed upstream of solenoid (not shown).
OLD DESIGN

## Slide 3
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
From
Tower
To
Pump
Sampling
Canister
Closed
Open
P
Open
Smart REA Sampler
Collection Mode
Identical systems for ‘up’ and ‘down’ sampling.
Need to minimize dead volume between solenoid and critical orifice.
Pressure gauge needs to work across wide range… something like <0.1 psi to 30 psi.
Manual valves must be leak tight in vacuum, and solenoid must seal under same conditions.
All tubing & fittings downstream of solenoid are stainless steel.
Filter needed upstream of solenoid (not shown).
Switching
3
To
Pump
3

## Slide 4
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
Sampling
Canister
Open
Open
P
Closed
Smart REA Sampler
Pressurization Mode
After sampling period, canister is filled to ~30 psi with UHP nitrogen.
Pressure gauge records pressure change so that the dilution factor is precisely known.
From
Regulated
Gas Cylinder
Open to Pump
3
To
Pump
3

## Slide 5
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
Sampling
Canister
Open
Open
P
Closed
Smart REA Sampler
Transfer Mode
Pressurized main canister is used to fill several smaller canisters/cartridges.
Cartridges are attached in parallel so that all are filled to same pressure.
Cartridges can be then transferred to laboratory for analysis.
Pressure gauge records pressure change so that the cartridge pressures are known.
Cartridge design still TBD.
To Smaller
Evacuated
Cartridges
Open to Pump
3
To
Pump
3

## Slide 6
P
Legend
Manual, High Integrity On/Off Valve
High Precision Critical Orifice
High Precision Vacuum/Pressure Gauge
Fast, Low Dead Space Three-Way Valve
Sampling
Canister
Open
Open
P
Open
Smart REA Sampler
Evacuation Mode
Pump must be able to evacuate system to near vacuum (~0.1 psi or below).
Solenoid also evacuated to minimize contamination.
To Suitable
Vacuum
Pump
Open to Pump
3
To
Pump
3